## Slot Parameters
| Application | HR/HRV/Soft adt green | SPO2 |  | Soft adt green |  | ADT
| Slot type | PPG | PPG | PPG | PPG | PPG | ADT
| LED color | Green | Red | IR | IR | NULL | IR
Slot parameters | SAMPLE_RATE_DIVIDER | 39(25Hz) | 39(25Hz) | 39(25Hz) | 39(25Hz) | 39(25Hz) | 199(5Hz)
| LED_SELECT | Select Green LED | Select Red LED | Select IR LED | Select IR LED | NULL | Select IR LED
| DRVx_CURRENT | 20mA(Total) | 40mA(Total) | 40mA(Total) | 10mA(Total) | 0mA(Total) | 10mA(Total)
| BG_LEVEL | 1 | 1 | 1 | 1 | 0 | 1
| RX0_PD_SEL | 0 | 0 | 0 | 0 | 0 | 0
| RX0_ASSIGN | HR/HRV/Soft adt green | SPO2 | SPO2 | Soft adt green | Soft adt green | ADT
| RX1_PD_SEL | 1 | 32 | 32 | 32 | 32 | 32
| RX1_ASSIGN | HR/HRV | - | - | - | - | -
| RX0_TIA_GAIN | 250K | 25K | 25K | 50K | 50K | 50K
| RX1_TIA_GAIN | 250K | 25K | 25K | 50K | 50K | 50K
| BG_CANCEL_ENABLE | 1 | 1 | 1 | 1 | 0 | 1
| RX0_BG_CANCEL | 0:32uA | 0:32uA | 0:32uA | 3:256uA | - | 3:256uA
| RX1_BG_CANCEL | 0:32uA | 0:32uA | 0:32uA | 3:256uA | - | 3:256uA
| DC_CANCEL_ENABLE | 0 | 0 | 0 | 0 | 0 | 0
| RX0_DC_CANCEL | 0 | 0 | 0 | 0 | 0 | 0
| RX1_DC_CANCEL | 0 | 0 | 0 | 0 | 0 | 0
| INT_TIME | 39.4us | 157.5us | 157.5us | 9.85us | 9.85us | 9.85us
| SYNC_OUTPUT_ENABLE | 0 | 0 | 0 | 0 | 0 | 0
| AGC_ENABLE | depend on hardware | depend on hardware | depend on hardware | 0:disable | 0:disable | 0:disable
| AGC_ADJ_UP_LIMIT (DRV code) | 255 (Depend on LED maximum load current) | 255 (Depend on LED maximum load current) | 255 (Depend on LED maximum load current) | - | - | -
| AGC_ADJ_DOWN_LIMIT (DRV code) | 16 | 16 | 16 | - | - | -
| AGC_SOURCE | Select the channel with the largest CTR | Select the channel with the largest CTR ; The distance from PD to red LED and IR LED must be equal | Select the channel the same with Red LED Slot ; The distance from PD to red LED and IR LED must be equal | - | - | - | For example, the multi-PD structure is shown in the figure above. PD selection rule of SPO2: Select the channel as AGC source with the largest CTR from PD0 or PD2 because they have the same distance to the red LED and IR LED; PD selection rules of HR: If you need to select two PDs from the four PDs, then select PD1 and PD3 because they are in the vertical arm direction; And select the channel with the largest CTR as AGC source from all the selected PDs.

## AGC parameters
| Application | HR/HRV/Soft adt green | SPO2
| Slot type | PPG | PPG | PPG
| LED color | Green | Red | IR
AGC | TRG_ADJUST_NUM | - | - | -
| LED_MAX_RAWDATA | - | - | -
| LED_MIN_RAWDATA | - | - | -
| LED_REF_UP | - | - | -
| LED_REF_DOWN | - | - | -
SOFT AGC | GAIN_ADJ_EN | 1 | 1 | 1
| BG_CANCEL_ADJ_EN | 0 | 0 | 0
| AMB_SAMPLE | 0 | 0 | 0
| AMB_SLOT_SEL | - | - | -
| GAIN_LIMIT_BG_32UA | 7 | 7 | 7
| GAIN_LIMIT_BG_64UA | 6 | 6 | 6
| GAIN_LIMIT_BG_128UA | 5 | 5 | 5
| GAIN_LIMIT_BG_256UA | 4 | 4 | 4
| AGC_TRIG_THD_H | 15518924 | 15938355 | 15938355
| AGC_TRIG_THD_L | 9646899 | 12582912 | 12582912
| AGC_RSTRN_THD_H | - | - | -
| AGC_RSTRN_THD_L | - | - | -