## Ver.Info (hidden)
Version | Note | Update Date | Author
0.1 | 首版 | 2016-10-27 | Leo
0.2 | 修订测量准确性场景 | 2016-11-02 | Leo
1 | 在准确性测试预置条件中去除“处于心率测量状态” | 2016-11-22 | Leo

## Test preparation
Test Equipments
1 | The test cellphone can be connected with the demo which to be tested.
2 | An APP that can record the measured output value of the demo to be tested is installed on the test cellphone.
3 | HR: Polar chest strap (H10/H7).
4 | SPO2: The oximeter of the top manufacturer (Masimo/Philips/Heal Force/Yuwell) is used as the secondary standard to output the reference results. It is recommended to use a medical grade monitor(such as Masimo) under permitting conditions.
5 | HRV：GOOD FRIEND_Single-lead Dynamic electrocardiogram instrument.
People Selection
1 | 10 people (or more than 10 people) were randomly selected to participate in each test scenario.
2 | The ratio of male to female is close to 1:1.
3 | The sample covered at least 1 obese person(BMI≥28.0) and 1 underweight person(BMI≤18.5).
Wearing Requirements of the watch/band to be tested
Watch/band：
1 | Wearing position at least 1 finger width away from the wristbone (2 finger width is recommended, which can be about 2.5cm away from the wristbone)
2 | After wearing the watch/band, no visible displacement of the watch/band during the arm swinging action.
GOOD FRIEND_Single-lead Dynamic electrocardiogram instrument
DL-191

## Version selection
Description of Differences
1 | The three versions of the test case have an inclusion relationship: Lite Version⊂Classic Version⊂Standard Version
Selection Principle
1 | According to the communication with the product customers, under the comprehensive consideration of at least include the existing acceptance scenarios of customers, the professionalism and depth of customers' cognitive and understanding of the relevant features. Choose one of the Lite version, classic version and standard version
2 | If the customer does not have standardized test cases or does not clear the test operation process, the lite version is recommended.
3 | If the customer has a reference test case or has certain test requirements but the specification is unclear and relatively simple, the classic version is recommended.
4 | If the customer has a clear understanding of the relevant features and principles, a complete and standardized set of acceptance cases, and clear access requirements for the system, the standard version is recommended.
5 | The test case provides two sets of test result evaluation criteria: absolute standard value evaluation and relative evaluation(compare with competitive products), which can be selected according to the customer's past usage criteria learnt from the communication with the customer.
The Usage of Document
1 | When sending test cases to customers, select one of the three versions.The version prompt in parentheses of the sheet name should be deleted, and only "Test Case" should be reserved the sheet name.
2 | In the Selected "Test Case" Sheet, select "Judgement Criteria-A" or "Judgement Criteria-B" from the test results. Delete the suffix letters from the reserved criteria column name and only the "criteria" should be reserved as column name.
3 | Delete "Revision Record" when release this document.
4 | Delete this sheet before release.

## Lite Version
Category | Item | Environment | Preparatory Conditions | Test Steps | Judgment Criteria-A | Judgment Criteria-B | Test Results | Interpretation of test results
Accuracy of HR detection
Rest
|  | 1.Demo 2.Polar chest strap (H10/H7) | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2. Rest in a sitting or lying position without data collection. 3. Start the demo and test for 5mins, the body remains still during the test. | 10 persons or more ±5BPM average of percentage statistical ≥95% ±10BPM average of percentage statistical ≥98% ±15BPM average of percentage statistical ≥100% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive products ±10BPM comparison of the average of percentage statistical and competitive products ±15BPM comparison of the average of percentage statistical and competitive products
Outdoor Jogging
|  | 1.Demo 2.Polar chest strap (H10/H7)） | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software 2. Environmental requirement: Outdoor flat pavement at the juction of sunlight and shadow. 3. 0:00-5:00 Jogging (5-8km/h) . 4. 5:00-10:00 Varied pace running (5-12km/h),it is recommended that HR>150bpm. 5. 10:00-15:00 Jogging (5-8km/h) . 6. 15:00-16:00 Rest. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Accuracy of SPO2 detection
Routine test
|  | 1.Demo 2.Oximeter | 1. SPO2 data recording software of Demos to be tested. 2. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testees sits and wears the demo, which is about 2.5cm away from the wrist joint with appropriate tightness. 2. The oximeter probe is clamped on the middle or index finger of the hand which wears the demo, and put hands on the table. 3. When the oximeter detects the SPO2 value and the value is stable, Start the test and last for 1min, the testees are required to keep silence during the test. 4. Repeat step 3 10 times, and readjust the wearing position of the demo before each test. | 10 persons or more RMSE≤3 | 10 persons or more Comparison of RMSE and competitive products
Accuracy of HRV detection
Routine test (PPG)
|  | 1.Demo 2.Sinle-lead Dynamic electrocardiogram instrument (GOOD FRIEND) | 1.Wear the Dynamic electrocardiogram instrument (GOOD FRIEND) 2.HRV data recording software of Demos to be tested 3.On line application of GOOD FRIEND 4. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testees wear the demo to be tested with appropriate tightness and wear the Dynamic electrocardiogram instrument (GOOD FRIEND), then put hands on the table and rest 1min. 2. Keep the sitting position and start the HRV test of demo and GOOD FRIEND for 10mins. 3. Exit the test when test ends. | 10 persons or more MAE≤10ms | 10 persons or more comparison of MAE and competitive products
Accuracy of wear detection （PPG）
Wear on/off detection
Normal wear
|  | 1.Demo | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo with appropriate tightness. 3. Remove the demo，then make the detection side of demo to air. 4.Repeat 2,3 steps 3 times. | 10 people 2. If correctly judged as wearing on status, record wear on once. 3.If correctly judged as wearing off status, record wear off once. 4. True wear on rate is 30/30，true wear off rate is 30/30 | 2~4. Count the true wear rate and true wear off rate of 10 people compared with competitive demo.
False wear off detection
Normal wear
|  | 1.Demo | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo with appropriate tightness. 3.Flip the wrist which wearing the demo 3 times and shaking hand 3 times. 4. Remove the demo，then make the detection side of demo to air. 5.Repeat 2-4 steps 3 times. | 10 people 4. If incorrectly judged as wearing off status, record false wear off once. 5. False wear off rate is 0/30 | 4~5.Count the false wear off rate of 10 people compared with competitive demo.
workout
|  | 1.Demo | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo with appropriate tightness. 3. Walk on a treadmill 2min（2-4km/h），then run 2min（6-10km/h） | 10 people 3. If incorrectly judged as wearing off status, record false wear off once; False wear off rate is 0/10 | 3. Count the false wear off rate of 10 people compared with competitive demo.
False wear on detection
Table
|  | 1.Demo 2.Table | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a table. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Printing paper
|  | 1.Demo 2.Printing paper | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a printing paper. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Mouse mat
|  | 1.Demo 2.Mouse | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a mouse mat. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
True wear off detection
Table
|  | 1.Demo 2.Table | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to the table (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Printing paper
|  | 1.Demo 2.Printing paper | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to a printing paper (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Mouse mat
|  | 1.Demo 2.Mouse mat | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to a mouse mat (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.

## Classic Version
Category | Item | Environment | Preparatory Conditions | Test Steps | Judgment Criteria-A | Judgment Criteria-B | Test Results | Interpretation of test results
Accuracy of HR detection
Rest
|  | 1.Demo 2.Polar chest strap (H10/H7) | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2. Rest in a sitting or lying position without data collection. 3. Start the demo and test for 5mins, the body remains still during the test. | 10 persons or more ±5BPM average of percentage statistical ≥95% ±10BPM average of percentage statistical ≥98% ±15BPM average of percentage statistical ≥100% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive products ±10BPM comparison of the average of percentage statistical and competitive products ±15BPM comparison of the average of percentage statistical and competitive products
Treadmill
|  | 1.Demo 2.Polar chest strap (H10/H7)） 3.Treadmill | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2. 0:00-0:45 Pace 5km/h. 3. 0:45-4:00 Pace 8-14km/h, Considering physical fitness, 8km/h is recommended . 4. 4:00-5:00 Pace 5km/h. 5. 5:00-6:00 Pace 8-14km/h, recommended 9-10km/h，faster than step 4. 6. 6:00-6:30 Pace 5km/h. 7. 6:30-7:30 Rest 1 min. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive products ±10BPM comparison of the average of percentage statistical and competitive products ±15BPM comparison of the average of percentage statistical and competitive products
Up and down stairs
|  | 1.Demo 2.Polar chest strap (H10/H7)） | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2.Go upstairs 20 floors (can also go upstairs 10 floors and go downstairs 10 floors), the testeess can control the speed of going up/down stairs according to their physical fitness. 3. Rest 1min. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Outdoor Walking
|  | 1.Demo 2.Polar chest strap (H10/H7)） | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2. Environmental requirement: Outdoor flat pavement at the juction of sunlight and shadow. 3. 0:00-10:00 Walk at normal speed. 4. 10:00-11:00 Rest. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Outdoor Jogging
|  | 1.Demo 2.Polar chest strap (H10/H7)） | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software 2. Environmental requirement: Outdoor flat pavement at the juction of sunlight and shadow. 3. 0:00-5:00 Jogging (5-8km/h) . 4. 5:00-10:00 Varied pace running (5-12km/h),it is recommended that HR>150bpm. 5. 10:00-15:00 Jogging (5-8km/h) . 6. 15:00-16:00 Rest. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Accuracy of SPO2 detection
Routine test
|  | 1.Demo 2.Oximeter | 1. SPO2 data recording software of Demos to be tested. 2. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testees sits and wears the demo, which is about 2.5cm away from the wrist joint with appropriate tightness. 2. The oximeter probe is clamped on the middle or index finger of the hand which wears the demo, and put hands on the table. 3. When the oximeter detects the SPO2 value and the value is stable, Start the test and last for 1min, the testees are required to keep silence during the test. 4. Repeat step 3 10 times, and readjust the wearing position of the demo before each test. | 10 persons or more RMSE≤3 | 10 persons or more Comparison of RMSE and competitive products
Hypoxia sit test
|  | 1.Demo 2.Oximeter 3. Plateau simulation equipment(MAG-20) | 1. SPO2 data recording software of Demos to be tested. 2. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testees sits and wears the demo, which is about 2.5cm away from the wrist joint with appropriate tightness. 2. The oximeter probe is clamped on the middle or index finger. 3. The testees sit on the chair with hands on both sides of his body. 4. Put on an oxygen mask 5. when the oximeter detects the SPO2 value and the value is stable, start the test，Stay at different hypoxic stages, each stage iss collected 5 times (1min each time). The SPO2 stages is as follows: a.SPO2 value is in the range of 90%-95% b.SPO2 value is in the range of 85%-90% c.SPO2 value is in the range of 80%-85% 6.Readjust the demo wearing position immediately before starting each test. In the process of adjusting the MAG-20 to collect the true value of SPO2 as recorded, the testees are required to keep silence. | 10 persons or more RMSE≤3 | 10 persons or more Comparison of RMSE and competitive products
Accuracy of HRV detection
Routine test (PPG)
|  | 1.Demo 2.Sinle-lead Dynamic electrocardiogram instrument (GOOD FRIEND) | 1.Wear the Dynamic electrocardiogram instrument (GOOD FRIEND) 2.HRV data recording software of Demos to be tested 3.On line application of GOOD FRIEND 4. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testees wear the demo to be tested with appropriate tightness and wear the Dynamic electrocardiogram instrument (GOOD FRIEND), then put hands on the table and rest 1min. 2. Keep the sitting position and start the HRV test of demo and GOOD FRIEND for 10mins. 3. Exit the test when test ends. | 10 persons or more MAE≤10ms | 10 persons or more comparison of MAE and competitive products
Accuracy of wear detection （PPG）
Wear on/off detection
Normal wear
|  | 1.Demo | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo with appropriate tightness. 3. Remove the demo，then make the detection side of demo to air. 4.Repeat 2,3 steps 3 times. | 10 people 2. If correctly judged as wearing on status, record wear on once. 3.If correctly judged as wearing off status, record wear off once. 4. True wear on rate is 30/30，true wear off rate is 30/30 | 2~4. Count the true wear rate and true wear off rate of 10 people compared with competitive demo.
False wear off detection
Normal wear
|  | 1.Demo | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo with appropriate tightness. 3.Flip the wrist which wearing the demo 3 times and shaking hand 3 times. 4. Remove the demo，then make the detection side of demo to air. 5.Repeat 2-4 steps 3 times. | 10 people 4. If incorrectly judged as wearing off status, record false wear off once. 5. False wear off rate is 0/30 | 4~5.Count the false wear off rate of 10 people compared with competitive demo.
workout
|  | 1.Demo | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo with appropriate tightness. 3. Walk on a treadmill 2min（2-4km/h），then run 2min（6-10km/h） | 10 people 3. If incorrectly judged as wearing off status, record false wear off once; False wear off rate is 0/10 | 3. Count the false wear off rate of 10 people compared with competitive demo.
False wear on detection
Table
|  | 1.Demo 2.Table | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a table. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Printing paper
|  | 1.Demo 2.Printing paper | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a printing paper. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Mouse mat
|  | 1.Demo 2.Mouse | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a mouse mat. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
True wear off detection
Table
|  | 1.Demo 2.Table | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to the table (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Printing paper
|  | 1.Demo 2.Printing paper | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to a printing paper (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Mouse mat
|  | 1.Demo 2.Mouse mat | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to a mouse mat (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.

## Standard Version
Category | Item | Environment | Preparatory Conditions | Test Steps | Judgment Criteria-A | Judgment Criteria-B | Test Results | Interpretation of test results
Accuracy of HR detection
Rest
|  | 1.Demo 2.Polar chest strap (H10/H7) | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2. Rest in a sitting or lying position without data collection. 3. Start the demo and test for 5mins, the body remains still during the test. | 10 persons or more ±5BPM average of percentage statistical ≥95% ±10BPM average of percentage statistical ≥98% ±15BPM average of percentage statistical ≥100% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive products ±10BPM comparison of the average of percentage statistical and competitive products ±15BPM comparison of the average of percentage statistical and competitive products
Life Scenarios
|  | 1.Demo 2.Polar chest strap (H10/H7) | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2. Sit on a chair and tie shoelaces 5 times. 3. Play mobile games for 1min while walking. 4. Sit on the chair in front of the desk and pick up and put down a pen 10 times. 5. Typing 3 minutes, recommended typing speed 40 words/min. 6. Rest 1 min. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Treadmill
|  | 1.Demo 2.Polar chest strap (H10/H7)） 3.Treadmill | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2. 0:00-0:45 Pace 5km/h. 3. 0:45-4:00 Pace 8-14km/h, Considering physical fitness, 8km/h is recommended . 4. 4:00-5:00 Pace 5km/h. 5. 5:00-6:00 Pace 8-14km/h, recommended 9-10km/h，faster than step 4. 6. 6:00-6:30 Pace 5km/h. 7. 6:30-7:30 Rest 1 min. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive products ±10BPM comparison of the average of percentage statistical and competitive products ±15BPM comparison of the average of percentage statistical and competitive products
Spinning
|  | 1.Demo 2.Polar chest strap (H10/H7)） 3.Spinning | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 3. 0:00-2:00 Easy to pedal (1-6 gears)，adjusted according to physical fitness. 4. 2:00-4:00 High intensity. speed >90 RPM, HR>120BPM. 5. 4:00-6:00 Same as step 3. 6. 6:00-8:00 High intensity. gear >9th and HR>150BPM. 7. 8:00-10:00 Same as step 3. 8. 10:00-11:00 Rest 1 min. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Elliptical machine
|  | 1.Demo 2.Polar chest strap (H10/H7)） 3.Elliptical trainer | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 3. 0:00-2:00 Low intensity (1-2 gears) 4. 2:00-4:00 moderate intensity (4-5 gears) 5. 4:00-6:00 Same as step 3. 6. 6:00-8:00 High intensity (6-8 gears) 7. 8:30-10:00 Same as step 3. 8. 10:00-11:00 Rest 1 min. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Steps test
|  | 1.Demo 2.Polar chest strap (H10/H7)） | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2 0:00-3:00 Male: 24 rounds per minute, 4 footbeats per round, 96 beats per minute in total; Female: 22 rounds per minute, 4 footbeats per round, 88 beats per minute in total. It is recommended to use Safe Metronome software to play the beat (step height: 40cm for male, 35cm for female). The specific movements of 4 footbeats are as follows: the first beat sounds, one foot is on the step; Second beat sounds, leg straight, another foot on the step and stand together; The third beat sounds,the foot that first on the step returns to the ground; The fourth beat sounds, the another foot also returns to the ground and turns into a ready posture. 3. 3:00-6:30 Sit down and Rest. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Up and down stairs
|  | 1.Demo 2.Polar chest strap (H10/H7)） | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2.Go upstairs 20 floors (can also go upstairs 10 floors and go downstairs 10 floors), the testeess can control the speed of going up/down stairs according to their physical fitness. 3. Rest 1min. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Outdoor Walking
|  | 1.Demo 2.Polar chest strap (H10/H7)） | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software. 2. Environmental requirement: Outdoor flat pavement at the juction of sunlight and shadow. 3. 0:00-10:00 Walk at normal speed. 4. 10:00-11:00 Rest. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Outdoor Jogging
|  | 1.Demo 2.Polar chest strap (H10/H7)） | 1. Wear the polar chest strap 2. HR data recording software of Demos to be tested and Polar 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testee wears the demo to be tested and Polar(H7/10) at the same time, and connects the demo and Polar(H7/10) via the HR data collection software 2. Environmental requirement: Outdoor flat pavement at the juction of sunlight and shadow. 3. 0:00-5:00 Jogging (5-8km/h) . 4. 5:00-10:00 Varied pace running (5-12km/h),it is recommended that HR>150bpm. 5. 10:00-15:00 Jogging (5-8km/h) . 6. 15:00-16:00 Rest. | 10 persons or more ±5BPM average of percentage statistical ≥80% ±10BPM average of percentage statistical ≥90% ±15BPM average of percentage statistical ≥95% | 10 persons or more ±5BPM comparison of the average of percentage statistical and competitive productss ±10BPM comparison of the average of percentage statistical and competitive productss ±15BPM comparison of the average of percentage statistical and competitive productss
Accuracy of SPO2 detection
Routine test
|  | 1.Demo 2.Oximeter | 1. SPO2 data recording software of Demos to be tested. 2. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testees sits and wears the demo, which is about 2.5cm away from the wrist joint with appropriate tightness. 2. The oximeter probe is clamped on the middle or index finger of the hand which wears the demo, and put hands on the table. 3. When the oximeter detects the SPO2 value and the value is stable, Start the test and last for 1min, the testees are required to keep silence during the test. 4. Repeat step 3 10 times, and readjust the wearing position of the demo before each test. | 10 persons or more RMSE≤3 | 10 persons or more Comparison of RMSE and competitive products
Hypoxia sit test
|  | 1.Demo 2.Oximeter 3. Plateau simulation equipment(MAG-20) | 1. SPO2 data recording software of Demos to be tested. 2. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testees sits and wears the demo, which is about 2.5cm away from the wrist joint with appropriate tightness. 2. The oximeter probe is clamped on the middle or index finger. 3. The testees sit on the chair with hands on both sides of his body. 4. Put on an oxygen mask 5. when the oximeter detects the SPO2 value and the value is stable, start the test，Stay at different hypoxic stages, each stage iss collected 5 times (1min each time). The SPO2 stages is as follows: a.SPO2 value is in the range of 90%-95% b.SPO2 value is in the range of 85%-90% c.SPO2 value is in the range of 80%-85% 6.Readjust the demo wearing position immediately before starting each test. In the process of adjusting the MAG-20 to collect the true value of SPO2 as recorded, the testees are required to keep silence. | 10 persons or more RMSE≤3 | 10 persons or more Comparison of RMSE and competitive products
Hypoxia lying test
|  | 1.Demo 2.Oximeter 3. Plateau simulation equipment(MAG-20) | 1. SPO2 data recording software of Demos to be tested. 2. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testees sits and wears the demo, which is about 2.5cm away from the wrist joint with appropriate tightness. 2. The oximeter probe is clamped on the middle or index finger. 3. The testees lies flat on the camp bed with hands on both sides of his body. 4. Put on an oxygen mask 5. when the oximeter detects the SPO2 value and the value is stable, start the test，Stay at different hypoxic stages, each stage iss collected 5 times (1min each time). The SPO2 stages is as follows: a.SPO2 value is in the range of 90%-95% b.SPO2 value is in the range of 85%-90% c.SPO2 value is in the range of 80%-85% 6.Readjust the demo wearing position immediately before starting each test. In the process of adjusting the MAG-20 to collect the true value of SPO2 as recorded, the testees are required to keep silence. | 10 persons or more RMSE≤3 | 10 persons or more Comparison of RMSE and competitive products
Low temperature test
|  | 1.Demo 2.Oximeter 3. Walk-in temperature laboratory | 1. SPO2 data recording software of Demos to be tested. 2. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testees sits quietly in the cryogenic laboratory (5℃) for 7mins, and wears the demo which is about 2.5cm away from the wrist joint with appropriate tightness. 2. The oximeter probe is clamped on the middle or index finger of the hand which wears the demo, and put hands on the table. 3. When oximeter detects the SPO2 value and the value is stable, Start the test and last for 1min, the testees are required to keep silence during the test. 4. Repeat step 3 10 times, and readjust the wearing position of the demo before each test. | 10 persons or more RMSE≤3 | 10 persons or more Comparison of RMSE and competitive products
Stance test
|  | 1.Demo 2.Oximeter | 1. SPO2 data recording software of Demos to be tested. 2. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. The testees stands quietly in the cryogenic laboratory (5℃) for 7mins, and wears the demo which is about 2.5cm away from the wrist joint with appropriate tightness. 2. The oximeter probe is clamped on the middle or index finger of the hand which wears the demo,Place hands flat on a platform at chest level 3. When oximeter detects the SPO2 value and the value is stable, Start the test and last for 1min, the testees are required to keep silence during the test. 4. Repeat step 3 10 times, and readjust the wearing position of the demo before each test. | 10 persons or more RMSE≤3 | 10 persons or more Comparison of RMSE and competitive products
Accuracy of HRV detection
Routine test (PPG)
|  | 1.Demo 2.Sinle-lead Dynamic electrocardiogram instrument (GOOD FRIEND) | 1.Wear the Dynamic electrocardiogram instrument (GOOD FRIEND) 2.HRV data recording software of Demos to be tested 3.On line application of GOOD FRIEND 4. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1.The testees wear the demo to be tested with appropriate tightness and wear the Dynamic electrocardiogram instrument (GOOD FRIEND), then put hands on the table and rest 1min. 2. Keep the sitting position and start the HRV test of demo and GOOD FRIEND for 10mins. 3. Exit the test when test ends. | 10 persons or more MAE≤10ms | 10 persons or more comparison of MAE and competitive products
Accuracy of wear detection （PPG）
Wear on/off detection
Normal wear
|  | 1.Demo | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo with appropriate tightness. 3. Remove the demo，then make the detection side of demo to air. 4.Repeat 2,3 steps 3 times. | 10 people 2. If correctly judged as wearing on status, record wear on once. 3.If correctly judged as wearing off status, record wear off once. 4. True wear on rate is 30/30，true wear off rate is 30/30 | 2~4. Count the true wear rate and true wear off rate of 10 people compared with competitive demo.
False wear off detection
Normal wear
|  | 1.Demo | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo with appropriate tightness. 3.Flip the wrist which wearing the demo 3 times and shaking hand 3 times. 4. Remove the demo，then make the detection side of demo to air. 5.Repeat 2-4 steps 3 times. | 10 people 4. If incorrectly judged as wearing off status, record false wear off once. 5. False wear off rate is 0/30 | 4~5.Count the false wear off rate of 10 people compared with competitive demo.
workout
|  | 1.Demo | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo with appropriate tightness. 3. Walk on a treadmill 2min（2-4km/h），then run 2min（6-10km/h） | 10 people 3. If incorrectly judged as wearing off status, record false wear off once; False wear off rate is 0/10 | 3. Count the false wear off rate of 10 people compared with competitive demo.
False wear on detection
Table
|  | 1.Demo 2.Table | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a table. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Printing paper
|  | 1.Demo 2.Printing paper | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a printing paper. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Mouse mat
|  | 1.Demo 2.Mouse | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a mouse mat. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Carton
|  | 1.Demo 2.Carton | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a carton. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Watchband
|  | 1.A demo with watchband | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Attach the detection side of the demo closely to watchband. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Roll paper
|  | 1.Demo 2.Roll paper | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo face to Roll paper. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
Quilt
|  | 1.Demo 2.Quilt | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo down to a quilt. 2. Start the test. 3. Repeat 1-2 steps 10 times. | 2. If incorrectly judged as wearing on status, record false wear on once. 3. False wear on rate ≤1/10 | 2~3. Count the false wear on rate of 10 tests compared with competitive demo.
True wear off detection
Table
|  | 1.Demo 2.Table | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to the table (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Printing paper
|  | 1.Demo 2.Printing paper | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to a printing paper (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Mouse mat
|  | 1.Demo 2.Mouse mat | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to a mouse mat (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Carton
|  | 1.Demo 2.Carton | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to a carton (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Watchband
|  | 1.A demo with watchband | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to the watchband (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Roll paper
|  | 1.Demo 2.Roll paper | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to a roll paper (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.
Quilt
|  | 1.Demo 2.Quilt | 1. Data recording software of Demos to be tested 3. (Optional) Wear competitive band/watch (The demo to be tested and competitive product wear one on each hand and test at the same time). | 1. Make the detection side of the demo to air, and start test. 2. Wear the demo. 3.Moving The detection side of the demo from the skin to the quilt (Unable to air the whole process). 4. Repeat 2-3steps 10 times. | 3. If correctly judged as wearing off status in 1min, record true wear off once, otherwise record unable to wear off. 4. True wear off rate ≥8/10 | 3~4. Count the true wear off rate of 10 tests compared with competitive demo.

## Terminology
Terminology | Explanation
Deviation ratio within ±5BPM | In all test data, the proportion of HR deviation between the test demo and standard equipment within ±5BPM (inclusive)
Deviation ratio within ±10BPM | In all test data, the proportion of HR deviation between the test demo and standard equipment within ±10BPM (inclusive)
Deviation ratio within ±15BPM | In all test data, the proportion of HR deviation between the test demo and standard equipment within ±15BPM (inclusive)
RMSE | Root Mean Squared Error, also known as the standard error, which is defined as , i=1, 2, 3,...n. In the limited number of measurements, the root mean square error is usually expressed by the following formula: √[∑(di^2)/n], where n is the number of measurements; di is the deviation between a set of measurement values and the true value
MAE | Mean Absolute Error is the average of absolute errors, which is defined as , i=1, 2, 3, ... n. In the limited number of measurements, the mean absolute error is usually expressed by the following formula: (∑¦di¦)/n, where n is the number of measurements; di is the deviation of a set of measurement values from the true value

## Revision Record
Version | Revision Date | Revision Content
1.0 | 2025-05-09 | The First Edition